· J1 : What kind of potting do we use?
Ans: Urethane
· J2 : Is Material 316L equivalent to 1,4435 din reference? Material type 316 is equivalent to what ?
Ans: (Yamatake will prepare the answer later)
· J3 : Flanges: Where are they produced?
Ans: Japan and China.

· J4 : JUMO needs mechanical drawing for the Hook connection and for the flange sealing
Ans: (Yamatake will prepare it later)

· J5 : Why stability is 0,1% for DP#2 and 0,2% for the other models?
Ans: (Yamatake will prepare the answer later)

· J6 : What is about the price for Kema approval?
Ans: about 16200 Euro (1Euro=\150)

· J7 : Looking the sample, JUMO has seen that the two filing screws are not protected at the bottom of the meter body. They need to be sealed
Ans: These are ball sealed.  Holes are bridged by silicon.

· J8 : Can JUMO have picture of our new transmitter. This is to have idea about their future design.
Ans: (Yamatake will prepare it later)

· J9 : What is the sensor resolution, the sensitivity, not the A/D converter but the sensor. i.e. the pressure difference for two adjacent digital values of the AD converter

Ans: (Yamatake will prepare the answer later)

· J10 : JUMO needs specification of our Hastelloy material
Ans: Alloy 276

· J11 : JUMO needs data sheet of the fill fluid oil (silicone and fluorine). 

·     Which fill oil is used for standard applications ? 

Ans: Both JUMO can have as a standard

·     Which fill oil is used for oxygen / chlorine application ? 

Ans: fluorine
·     Does Yamatake offer a special oil for food/pharma applications , with FDA approval ?
(For those special customers JUMO is using "ENERGOL WM 2" from supplier BP)
Ans: Yamatake can make a MB for another fill fluid by request base.  The dimension of MB which is added on structure will be different than standard MB.

· J12 : JUMO needs EEPROM type and mapping
Ans: Need to discuss by both engineers

· J13 : JUMO needs configuration of the A/D converter registers
Ans: Need to discuss by both engineers


· J14 : All sensors can be used for explosion safe transmitters, i.e. intrinsically safe and flame proof transmitters, and the required test procedures are performed during manufacturing by Yamatake. Is this correct?

Ans: Yes

· J15 : What is the exact meaning of accuracy in Yamatake table: Is it only "nonlinearity" or   "nonlinearity + hysteresis + repeatability" ?
Ans: nonlinearity + hysteresis + repeatability
· J16 : If possible, please send JUMO an offer for the flanges of your supplier. Annual quantity 2000 / 4000 / 6000 pcs.
Ans: (Yamatake will prepare the answer later)

· J17 : Please complete the price list for the other types of DP sensors (different material, fill oil, etc,)
Ans: (Yamatake will prepare the answer later)

