	Item
	Change to Software / Hardware
	Completed Date
	Details / MEETING NOTES for 8/20/2006 @ Yamatake

	1xx
	COM-PORT Property
	Completed 3/2006 V4.81
	Increase property up to port-4 to up to port 8 or 10

	2xx
	Analysis report
	Completed 3/2006 V4.80
	Should be not written as other value or blank but should be written as the value which was set, for example 14.6 psi

	3xx
	Printer format
	Completed 4/2005 Excel
	LETTER format instead of A4 (Will be included on Disc)

	4x
	Compressibility added to printed report
	Completed 4/2005 Excel
	Add compressibility to printed report (Will be included on disc)

	5xx
	Multi-stream software release
	Completed 8/2005
	Ability to run up to 4 sample streams

	6x
	Calibration report
	Next release
	(FIRMWARE) Add average result of calibration gas in HGM calibration report screen.  Place where (last analyzed results) are at present. Also add Btu, SG, total un-normalized.  Add ability to print this screen
MEETING NOTES:#6, OK, sending report for our review (Cosa)

	7x
	Time display in the status when it is running 
	Next release
	(SOFTWARE) Time status is displayed from 0 to 300 seconds (new Box below GC picture)  along with time scale on chromatogram report and on chromatogram display screen.  Add sliding pointer to show time on chromatogram display screen

	8x
	Remove all display of Peak Area
	Next release
	(SOFTWARE) PA is not used. Therefore it is better to remove it from all screens and reports

	9x
	Chromatogram report
	Next release
	(SOFTWARE) Remove PA (peak area) columns and add total of un-normalized below un-normalized column of mol%, response factor of each component to replace peak area column

	10x
	Show Stream #  & Name on Analysis Report & HGM
	Next Year
	(SOFTWARE) Printed Report does not automatically designate stream number from stream switching SW.  Add stream number and name to printed analysis report, chromatogram report and on HGM screen.

	11x
	Add ability to print out all screens 
	Next Year
	(SOFTWARE) Some screens are not printable Calibration Record, Zoom chromatogram and calibration screen as examples.

	12x
	Speed up data processing at the end of run
	Next Year
	(SOFTWARE) Make HFA drivers run in the background to speed up software. When the GC is making chromatogram this takes time and you can not access any other feature during this process.  Could chromatogram be turned on and off to make data available faster? Basically speed up calculation and report generation process. If not fast enough, have chromatogram on demand to stop/start. 
MEETING NOTE:#12, only driver running in background. Chromatogram on/off not needed for speed improvement.

	13x
	Combine all software into one package.
	Next Year
	(SOFTWARE) Customers request all software to be combined into one package. HGM,HMV,HDS and SS.  Rename each tab so that the function of the particular tab is better described.
MEETING NOTE:#13, problematic

	14x
	Physical constant of the GPA table configurable with software
	Next Year
	(FIRMWARE) Presently only configurable thropugh firmware.  Add GPA table so that the factors can be updated in the future through software instead of firmware.

	15x
	Halt function
	Next Year
	(FIRMWARE/SOFTWARE) Add ability to halt the analyzer other than power down.
MEETING NOTES:#15, OK, add RUN(green) HALT(red)

	16
	Deg F and Psi instead of Deg C and Kpa
	Next release
	(FIRMWARE) Controller display needs to show deg F and Psig - software/reports already are correct 



	17x
	Password
	Next Release
	Make the default passwords (COSA)
MEETING NOTE:#17, OK, blank , no password #18, OK, but will confirm

	18x
	Synchronize 
	Next Year
	When synchronizing upon changes to stream switching software (multi-stream).  Give ability to minimize screen.  Add warnings to buttons that are more easily noticed by user. 

	19x
	Minimize Main Menu
	Next Release
	Add ability to minimize HGM main menu screen.

	20x
	Calibration screen
	Next Release
	Check boxes should be checked if all mol% have a value,  If mol% = 0 Then check box should not be checked.  Need warning if criteria is not met before closing.  Need warning when attempting to close calibration screen before all data has been downloaded to HGC.

MEETING NOTE:#20, warning window message: Component value entered but check box not selected (will call out values not checked in warning window). Same for Checked box with 0 (zero) entered, message: Component box selected but no component valve entered or zero value entered. Remove check or enter value. PRINT REPORT, not Print Screen

	21
	(not openly discussed) Modbus for C6+ Split  
	Next Release
	Add 8,9,10,11 because all modern  Flow Computers with the  Daniel Modbus load are expecting 8 instead of 108.


