[image: image12.png]



[image: image1.jpg]i \\\\\\
\\ 11

\\\\\
\\\\\ Q, I
Water

\)

2007-337999/001 001 g "
Tohde B/BL DA
Fan TS

D: 1kgh

T:20°C

V: 1 mPas R

p: 1,013 bar





[image: image11.png]MECE®N

FLOW-CONTROL-SYSTEMS




Operating instructions
HART® communication
[image: image2.jpg]WE’VE GOT

ABILITY




All rights reserved. It is prohibited to reproduce this documentation, or any part thereof, without the prior written authorisation of MECON Flow Control Systems GmbH.

Subject to change without notice.

Copyright 2011 by

MECON Flow Control Systems GmbH - Röntgenstraße 105 - 50169 Kerpen
Contents
1  Safety instructions
5

1.1
Intended use
5

1.2
Certifications
6
2  Device description
8

2.1
Scope of delivery
8

2.2
Description code
11

3  Installation
15

3.1
Notes on installation
15

3.2
Storage
15
4  Electrical connections
18

4.1
Safety instructions
18

4.2
Electrical connections
18

4.2.1
Current output (MEM)
19

4.2.2
Current output with digital outputs (MEM)
19
5  Start-up
22

5.1
Device with current output
22

5.2
Device with additional limit switches
23


5.3
Device with additional limit switch and pulse output
23

6  Service
24

6.1
Maintenance
24

6.2
Retrofitting transducer MEM
24

6.3
Returning the device to the manufacturer
26

6.4
Disposal
26
7  Technical data
27

7.1
Operating principle
27

7.2
Technical data
28
1  Safety instructions
1.1
General information
	 
	Important!
This manual is solely a supplement to the document »Operating in​struc​tions RE250«. 

To prevent any injury to the user or damage to the device it is essen​tial that you read the safety instructions in the »Operating instruc​tions RE250« and observe it unconditionally.

The manufacturer is not liable for any damage resulting from improper or unintended use of these devices.


1.2
Certifications
CE marking

[image: image3.jpg]



The manufacturer certifies for the device RE 250 with integrated transducer MEM the fulfilment of all statutory require​ments of the following EC directives by applying the CE marking:
•
Pressure equipment directive 97/23/EC 

•
Low voltage directive 2006/95/EC



•
EMC-directive 89/336/EEC
•
NAMUR recommendation NE21

2  Device description
2.1
Scope of delivery
[image: image4.jpg]



Fig 2.1 
Scope of delivery
	(
	Information!

Please check the delivery for completeness using the packing list. 


2.2
Description code
The specific output configuration of the device can be taken from the description code:

[image: image5.jpg]Ty 5 @ A AT @ 1) @ 13 @9 1>




	

	F A
	With HART protocol, 4 - 20mA

	F L
	With HART protocol, 4 - 20mA, 2 inductive contacts SJ 3,5N

	F J
	With HART protocol, 1 inductive contact, 1 pulse output

	
	


3  Installation

3.1
Notes on installation

	(
	Information!

All instruments are carefully checked for functional capability before distribution. On receipt of the device please carry out a visual inspec​tion of the outer packing for damage or improper handling.
Please contact the carrier and your responsible sales field service if you should disco​ver any defects. In such cases a description of the defect, the type and the serial number of the device is needed.

	(
	Information!

Unpack the equipment with care to prevent damage.

	(
	Information!

All instruments are carefully checked for order conformity.

Please check the delivery for completeness using the packing list. Please examine the flowmeter nameplate to verify that the device was built according to your order. Particularly check devices with electrical components for the correct supply voltage.

	(
	Information!

This manual is solely a supplement to the document »Operating instruct​tions RE250«. To prevent any injury to the user or damage to the device it is essential that you read the installation notes in the »Operating in​structions RE250« and observe it unconditionally.


3.2
Storage

●
Store the device in a dry and dust-free place.

●
Keep away from direct sun and heat.

●
Avoid external load to the device.

●
The storage temperature range for standard devices with electrical components is about -40 … +70 °C / -40 … +158 °F. 

4  Electrical connections
4.1
Safety instructions
	 
	Warning!

All work on the electrical connections may only be carried out with the power disconnected. 

It is essential to observe the voltage data on the nameplate.

	 
	Warning!

Observe all occupational safety regulations unconditionally.
Any work done on the electrical components of the device may only be carried out by a specialist.

	
	Caution!

After installation and any maintenance work on the device, the operator has to check and ensure the specified protection class again.


4.2
Electrical connections

4.2.1
Current output (MEM)
[image: image6.jpg]Hazardous area

Angular sensor

type MEM

_Aa

1

Transmitter power supply

-+

Safe area

+H-0

4-20mA
2

-0

Supply current circuit

Li = 240pH Ui = 30V

Ci = 16nF li = 150 mA
Ta = 70°C Pi = 1000 mW
EExiallCT6

BVS 07 ATEXE 033

E

/

E

Qutput 4 -20 mA

Auxiliary power suppIyZ>

HART-Modem

Rmin
2500




Fig 4.1
Connection diagram for transducer MEM with current output

4.2.2
Current output with digital outputs (MEM)

[image: image7.jpg]Hazardous area

Supply current circuit

Safe area

BVS 07 ATEXE 033

|
|
|
Li = 240pH Ui = 30V :
Ci = 16nF  li = 150 mA Transmitter power supply
Ta = 70°C  Pi = 1000 mwW |
EEx ia IIC T6 | + E
Angular sensor BVS 07 ATEXE 033 |
type MEM : E 4 -20 mA
f ol + _01 : = /
4-20mA | i
o2 |
| Auxiliary power supply
3 |
MAX :4 [
# 1O ! Isolation amplifier
5 | N
# ——
MIN _06 | ] \ MAX limit value
# O | # bbb
| o
Contacts according to : | #
EN 60947-5-6 ~ EinamiouipuG | |8 2 MIN limit value
(NAMUR) Li = 4pH U = 30V )
Ci = 16nF  li = 20mA |
Ta = 70°C Pi = 100mw |
EEx ia IIC T6 | Auxiliary power supply
|
|




Fig 4.2
Connection diagram for transducer MEM with 2 limit switches
[image: image8.jpg]Hazardous area

Supply current circuit

Safe area

Transmitter power supply

E

\

i

4-20 mA

4 Auxiliary power supply

Isolation amplifier

=

BVS 07 ATEX E 033

AN

I
I
I
Li = 240uH Ui = 30V '
Ci = 16nF li = 150 mA |
Ta = 70°C  Pi = 1000mw |
EExia llC T6 | +
Angular sensor BVS 07 ATEX E 033 |
type MEM |
I
ﬁ o+ _01 | -
I
4-20 mA |
2
E [© :
3 I
TULL —O I
# _04 |
5 | -
MIN —O I
# o8 ! :
| %
Contacts according to : [
Binary outputs )
E’\(IN6A?I\9/I?J7I§)S_6 Li = 4pH Ui = 30V |
Ci = 16nF i = 20mA |
Ta = 70°C Pi = 100mw |
EExia llC T6 :
|

#

#/

Pulse output

MIN/MAX limit value

4 Auxiliary power supply





Fig 4.3
Connection diagram for transducer MEM with limit switch and pulse output
5  Start-up

5.1
Device with current output (MEM)
The magneto-electrical measuring transducer (MEM) is completely factory-set when it is delivered to the customer.

After applying the supply voltage to the device, initially the current output will be about 3.5 mA to 4 mA for a few seconds. After that a current corresponding to the pointer deflection will flow.
	(
	Important!

Due to the influence of the float magnet, the MEM transducer will only output the correct current if the pointer position is caused by the float. Turning the pointer manually will cause incorrect values, but it’s suitable for testing.


5.2
Device with additional limit switches

The limit switches are integrated into the angle transducer MEM and the design is according to NAMUR recommendation NE21. 
The switching points can be set independently over the entire measuring range by a HART® command. For details please refer to chapter ”xx”.
	(
	Important!
Unless otherwise desired by the customer, the factory set value for the switching point is about 20% / 80% of URV and the switching function is active state »OPEN«.


5.3
Device with additional limit switch and pulse output

One of the integrated limit switches (design is according to NAMUR recommendation NE21) can be configured as a pulse output. The pulse unit is identical to the unit of the measured value.

The switching point of the second switch can be set independently over the entire measuring range by a HART® command. For details please refer to chapter ”xx”.
	(
	Important!

Unless otherwise desired by the customer, the factory set value for the switching point is about 40% of URV and the switching function is active state »OPEN«.


6  Operation
6.1
Parameter setting using HART protocol
The parameter of the MEM can be set using a hand-held terminal or a PC using a suitable HART®-modem.

The following parameters can be set:

	Main menu
	Level 2
	Level 3
	Level 4
	Description

	Device setup
	Process variables
	Flow rate
	
	Actual flow rate

	
	
	Totalizer
	
	Actual totalizer value

	
	Diag/service
	Test device
	Self test
	Self test parameters

	
	
	
	Status
	Error indication

	
	
	Simulation
	Default values
	Values for simulation

	
	
	
	Start simulation
	Start/stop simulation

	
	
	Calibration
	
	Calibration parameter

	
	
	DA trim
	
	Current output trim

	
	
	Reset device
	
	Warm start of device

	
	Basic setup
	TAG
	
	TAG number

	
	
	PV unit
	
	Unit of measured value

	
	
	Range values
	PV LRV
	’0’

	
	
	
	PV URV
	Upper range value

	
	
	
	PV unit
	Unit of measured value

	
	
	
	PV LSL
	Lower limit for MR

	
	
	
	PV USL
	Upper limit for MR

	
	
	Device information
	Model
	’ES’ 

	
	
	
	Device ID
	Unique device ID

	
	
	
	TAG
	TAG number

	
	
	
	Date
	Calibration date

	
	
	
	Write protect 
	’NO’

	
	
	
	Descriptor
	Description of device

	
	
	
	Message
	Brief message

	
	
	
	Revision #
	Revision number

	
	
	Xfer function
	
	’Linear’

	
	
	PV damp
	
	Time constant

	
	Detailed setup
	Totalizer
	
	Totalizer function on/off

	
	
	Low flow cutoff
	
	Cutoff value in %

	
	
	Namur
	MAX
	MAX value in % of URV

	
	
	
	MIN
	MIN value in % of URV

	
	
	
	Switch function
	N1-N2, MIN-MAX, …

	
	
	Switch
	
	Active state open/closed

	PV
	
	
	
	Measured value

	PV AO
	
	
	
	Anaolg output (corresp.)

	LRV
	
	
	
	Lower range value

	URV
	
	
	
	Upper range value


7  Service

7.1
Maintenance
The devices were built within the scope of low maintenance but periodically the flow​meters should be inspected for signs of corrosion, mechanical wear as well as da​mage to the fitting and the display unit.
We advice to carry out inspections at least once a year.

For a detailed inspection and cleaning the device must be removed from the piping.

	 
	Caution!

Appro​priate safety precautions have to be taken when removing the device. Always use new gaskets when reinstalling the device in the piping. 




7.2
Retrofitting transducer MEM
Devices with measuring ring

For devices with standard measuring ranges starting from 5-50 l/h (gases: 0.15 – 1.5 m³/h) the float damping can be replaced by the customer:

●
Remove the device from the piping.

●
Fix the device, ensuring that the fitting will not be damaged.

●
The damping cylinder (  is fixed by re​taining clamps (. Bend up these two clamps carefully using a suitable tool.
●
Move the damping cylinder a little bit to the bottom end of the fitting, make a quarter turn and take the cylinder out of the fitting.

●
Insert the new damping cylinder into the fitting from the top, slide it over the damping part of the float and arrange it into the retaining clamps.

●
Fix the cylinder by bending the retaining clamps together.

●
Fit the device back into the piping.
7.3
Returning the device to the manufacturer
This device has been carefully manufactured and tested. Should you nevertheless need to return a device to MECON GmbH please observe the following points:
	 
	Caution!

According to the actual waste disposal directives, the owner/customer is responsible for the waste management of hazardous and toxic waste.
For reasons of environmental protection and safeguarding the health and safety of our personnel all devices sent to MECON GmbH to be repaired must be free of toxic and hazardous substances. This also applies to cavities of the devices. If necessary the customer is kindly requested to neutralize or rinse the devices before returning to MECON.
The customer has to confirm this by filling in an appropriate form which is available for download on the MECON website: 

www.mecon.de/en/declaration/contamination.pdf



	 
	Caution!

In the case of returning devices, despite these requirements, which contain hazardous and toxic substances, Mecon GmbH is entitled to dispose these substances at the cost of the customer without any further inquiries.




7.4
Disposal
	 
	Caution!

Disposal of the devices has to be carried out in accordance with relevant legislation in your country.




8  Technical data

8.1
Operating principle
Like other devices of this series the variable area flowmeter RE250 operates on the principle of flotation:
The measuring unit consists of a metal tube with a measuring ring in which a float can move up and down. The movement of the float is transmitted via a magnet to a subsequent magnet which is mounted on the pointer axis.

The position of this magnet is detected by a Hall-sensor inside the integrated angle transducer MEM (sensor signals A, B) and generates a resulting output current of 4...20mA. The signal is linearised with a maximum of 16 sampling points in the process. 
[image: image9.jpg]SN NN N N N N N N N N N N N N N
\

= | o]
\

\
NN N N N N N N N N N N N N N N N Y N N N N N N N N N N

NN N N N N N N N N N N N N N N N NN

C N
yJ

BN N N N N N N N





Fig. 7.1
Operating principle of the RE250

8.2
Technical data
General data
	Range of application
	Flow measurement of liquids and gases

	Measuring principle
	Flotation / Float measuring

	Orientation
	Vertical – flow direction from bottom to top


Measuring accuracy*
	Directive
	VDI / VDE 3513, sheet 2 (qG = 50%)

	Liquids
	G 1.6 (add. 0,2 %** of URV for MEM)

	Gases
	G 2.0 (add. 0,2 %** of URV for MEM)

	Reproducibility
	0,5 % of URV (add. 0,1 % of URV for MEM)


*   A variation of the operating temperature to the considered temperature for the calibration process, will lead to a corresponding inherent error

** For ambient temperatures < -20 °C the measuring error will increase
Temperatures
	Device version
	Ambient temperature*
	Storage temperature

	
	[°C]
	[°F]
	[°C]
	[°F]

	with 4 … 20 mA + HART® protocol
	-40 …+70
	-40 …+158
	-40 …+70
	-40 …+158

	with 4 … 20 mA + HART® + binary outputs
	-40 …+70
	-40 …+158
	-40 …+70
	-40 …+158

	 
	* Danger!

For applications in hazardous areas it is mandatory that the temperature class of the type examination certificate (protection type) will be regar​ded additionally.

For devices used in hazardous areas, additional safety notes apply; please refer to the Ex documentation!



RE250 with current output 4 …20 mA and HART® protocol
	Cable gland
	M20 x1.5

	Terminal connection
	2.5 mm² 

	Auxiliary supply voltage UB
	14 V … 30 V DC

	Measuring signal
	4 … 20 mA = 0 … 100% flow value in 2-wire technology

	Power supply influence
	< 0.1 % of MV

	Max. external load RB
	680 Ω (30 V)  
RB = (UB - 14V) / 22 mA

	Min. external load RB
	250 Ω 

	Temperature influence
	< 10 µA / K

	Min. power supply
	19.5V DC

	HART protocol revision
	5.0

	Device revision
	1


RE250 with current output 4 …20 mA and HART® protocol + binary outputs
Supplements to technical data for RE250 with current output 4 … 20 mA:

	Binary outputs

	Switching function
	NAMUR (EN 60947-5-6:2000) 

	Max. voltage
	30 V DC

	Max. power consumption
	100 mW

	Current on state OPEN
	typ. 0.4 mA

	Current on state CLOSED
	typ. 4 mA

	Pulse output

	Max. Rate
	10 Hz

	Pulse width
	approx. 50 ms


	Mecon GmbH

Röntgenstraße 105

D 50169 Kerpen

Tel.:
+49 (0)2237 600 06 -0

Fax.:
+49 (0)2237 600 06 – 20

Email: 
customerservice@mecon.de
www.mecon.de
	[image: image10.jpg]TUVRheinland®

CERT

ISO 9001

A

-4 7
TUVRheinland

v o
2N







  RE 250  Variable Area Flowmeter














1  Float with damping


2  Upper guide bracket


3  Self-locking nut


4  Lower guide bracket


5  Retaining clamp


6  Damping cylinder





1	Flowmeter RE250 (with integrated transducer MEM)


2	Operating Instructions


3	Certificates (optional)


4	Software

















© MECON GmbH 04/2011  OI_RE250_HART_EN


